- Brief introduction

The Urethral Catheters (Latex) are flexible tubes placed in the body to drain and collect urine from the bladder.

The Urethral Catheters are flexible tubes that are passed through the urethra during urinary catheterization and into the bladder to drain urine, or for inserting fluids into the bladder.

Urethral Catheter is used in departments of urology, internal medicine, surgery, obstetrics, and gynecology for drainage of urine and medication. It is also used for patients suffering form moving with difficulty or being completely bed-ridden. It is easy to use, reliable in performance, and irritation free.

The Urethral Catheters come in four types, it includes Latex Foley Catheter, Tiemann Urethral Cathete, Latex Urethral Catheters,Red Urethral Catheter.

The Urethral Catheters are made from natural latex , consists of shaft, drainage funnel, inflation funnel, irrigation funnel (if present), balloon (if present)，valve (if present) Activation Rod(if present) ,Bellows (if present).and Connector(if present). The Urethral Catheter has the following features:

- Silicone coated

- Color-coded for visualization of size

- Alternative types are available

- Sterile by EO

- For single use only

- Hard & soft valves are available

The Urethral Catheters are used during the following situations:

- On patients who are anesthesized or sedated for surgery or other medical care

- On comatose patients

- On some incontinent patients

- On patients whose prostate is enlarged to the point that urine flow from the bladder is cut off. The catheter is kept in until the problem is resolved.

- On patients with acute urinary retention.

- On patients who are unable due to paralysis or physical injury to use either standard toilet facilities or urinals.

- Following urethral surgeries

- Sometimes before Furosemide administration

Table of Variants

| Type | Sub-type | Size |
| --- | --- | --- |
| Latex Foley Catheter | Pediatric, 2-way | 6Fr, 8 Fr, 10 Fr |
|  | Female, 2-way | 12Fr, 14 Fr, 16 Fr, 18 Fr, 20 Fr, 22 Fr, 24 Fr, 26 Fr |
|  | Male(standard), 2-way | 12Fr, 14 Fr, 16 Fr, 18 Fr, 20 Fr, 22 Fr, 24 Fr, 26 Fr, 28 Fr, 30 Fr |
|  | 3-way | 14 Fr, 16 Fr, 18 Fr, 20 Fr, 22 Fr, 24 Fr, 26 Fr |
|  | 4-way | 18 Fr, 20 Fr, 22 Fr, 24 Fr |
|  | 5-way | 18 Fr, 20 Fr, 22 Fr, 24 Fr |
| Tiemann Urethral Catheter | 2-way | 12Fr, 14 Fr, 16 Fr, 18 Fr, 20 Fr, 22 Fr, 24 Fr, 26 Fr, 28 Fr, 30 Fr |
|  | 1-way | 12Fr, 14 Fr, 16 Fr, 18 Fr, 20 Fr, 22 Fr, 24 Fr, 26 Fr, 28 Fr, 30 Fr |
| Red Urethral Catheter | 1-way | 6Fr, 8 Fr, 10 Fr, 12Fr, 14 Fr, 16 Fr, 18 Fr, 20 Fr, 22 Fr, 24 Fr |
| Latex Urethral Catheter | 1-way | 6Fr, 8 Fr, 10 Fr, 12Fr, 14 Fr, 16 Fr, 18 Fr, 20 Fr, 22 Fr, 24 Fr |

2. The purpose and functions of the product and its variants and their design

Intended Use: The Urethral Catheter is used in departments of urology, internal medicine, surgery, obstetrics, and gynecology for drainage of urine and medication. It is also used for patients suffering form moving with difficulty or being completely bed-ridden. The Urethral Catheters are passed through the urethra during urinary catheterization and into the bladder to drain urine, or for inserting fluids into the bladder.

To achieve above-mentioned purpose, the product should have the following functions: draiage of urine and/or inserting fluids into the bladder for medication.

Drainage of urine or medication:

The Urethral Catheter is made of natural latex which has appropriate rigidity, it resists to kink and soft enough for urinary catheterization. After inserting, the drainage lumen of the Urethral Catheter shall be able to drain the urine or remove the fluids into the bladder.

3. Classification according to Annex IX of Council Directive 93/42/EEC

The Urethral Catheter is invasive device. It is intended for long-term use. According to rule 5 of Annex IX, III of Council Directive 93/42/EEC (2007), the Urethral Catheter belongs to class Ⅱb medical device.

4. Main technical parameters of the product

4.1 Material

Material for Urethral Catheter as follows:

| No. | Component | Material used | Spec/Code | Supplier |
| --- | --- | --- | --- | --- |
| 1 | Shaft | Natural latex | VCT,PVH5003 | GETAHINDUS (M) SDN BHD |
| 2 | Funnels | Natural latex | VCT,PVH5003 | MMG polymer Company Limited |
| 3 | Balloon | Natural latex | VCT,PVH5003 | GETAHINDUS (M) SDN BHD |
| 4 | Sleeve | Natural latex | VCT,PVH5003 | MMG polymer Company Limited |
| 5 | Soft valve | Natural rubber | / | Zhanjiang fengfan |
| 6 | Hard valve | PVC, PP, ABS, Silicone and stainless steel | RE2-010/344R/757/—/— | Jiangxi zongjie |
| 7 | Coating | Silicone | / | Shanghai Xilikang |
| 8 | Primary Package | Inner pouch | Coated paper+film | coated paper 60g + PET/PE | Quzhou Zexiao Medical Packaging Material Co., Ltd |

4.2 Physical requirements

4.2.1 Measures

Table 1: Dimensions of Latex Foley Catheter, Pediatric, 2-way

| Size(Fr/Ch) | O.D, ±0.3mm | Overall length, mm | Flow rate, ml/min |
| --- | --- | --- | --- |
| 6 | 2.0 | 270±10 | ≥10 |
| 8 | 2.7 | 270±10 | ≥15 |
| 10 | 3.3 | 270±10 | ≥30 |

Table 2: Dimensions of Latex Foley Catheter, Female, 2-way

| Size(Fr/Ch) | O.D, ±0.3mm | Overall length, mm | Flow rate, ml/min |
| --- | --- | --- | --- |
| 12 | 4.0 | 255-270 | ≥50 |
| 14 | 4.7 | 255-270 | ≥70 |
| 16 | 5.3 | 255-270 | ≥100 |
| 18 | 6.0 | 255-270 | ≥100 |
| 20 | 6.7 | 255-270 | ≥100 |
| 22 | 7.3 | 255-270 | ≥100 |
| 24 | 8.0 | 255-270 | ≥100 |
| 26 | 8.7 | 255-270 | ≥100 |

Table 3: Dimensions of Latex Foley Catheter, Male(standard), 2-way

| Size(Fr/Ch) | O.D, ±0.3mm | Overall length, mm | Flow rate, ml/min |
| --- | --- | --- | --- |
| 12 | 4.0 | 400±10 | ≥50 |
| 14 | 4.7 | 400±10 | ≥70 |
| 16 | 5.3 | 400±10 | ≥100 |
| 18 | 6.0 | 400±10 | ≥100 |
| 20 | 6.7 | 400±10 | ≥100 |
| 22 | 7.3 | 400±10 | ≥100 |
| 24 | 8.0 | 400±10 | ≥100 |
| 26 | 8.7 | 400±10 | ≥100 |
| 28 | 9.3 | 400±10 | ≥100 |
| 30 | 10.0 | 400±10 | ≥100 |

Table 4: Dimensions of Latex Foley Catheter, 3-way

| Size(Fr/Ch) | O.D, ±0.3mm | Overall length, mm | Flow rate, ml/min |
| --- | --- | --- | --- |
|  |  |  | Drainage lumen | Irrigation lumen |
| 14 | 4.7 | 400±10 | ≥70 | ≥25 |
| 16 | 5.3 | 400±10 | ≥100 | ≥25 |
| 18 | 6.0 | 400±10 | ≥100 | ≥25 |
| 20 | 6.7 | 400±10 | ≥100 | ≥25 |
| 22 | 7.3 | 400±10 | ≥100 | ≥30 |
| 24 | 8.0 | 400±10 | ≥100 | ≥30 |
| 26 | 8.7 | 400±10 | ≥100 | ≥30 |

Table 5: Dimensions of Latex Foley Catheter, 4-way

| Size(Fr/Ch) | O.D, ±0.3mm | Overall length, mm | Flow rate, ml/min |
| --- | --- | --- | --- |
|  |  |  | Drainage lumen | Irrigation lumen |
| 18 | 6.0 | 400±10 | ≥100 | ≥25 |
| 20 | 6.7 | 400±10 | ≥100 | ≥25 |
| 22 | 7.3 | 400±10 | ≥100 | ≥30 |
| 24 | 8.0 | 400±10 | ≥100 | ≥30 |

Table 6: Dimensions of Latex Foley Catheter, 5-way

| Size(Fr/Ch) | O.D, ±0.3mm | Overall length, mm | Flow rate, ml/min |
| --- | --- | --- | --- |
|  |  |  | Drainage lumen | Irrigation lumen |
| 18 | 6.0 | 400±10 | ≥100 | ≥25 |
| 20 | 6.7 | 400±10 | ≥100 | ≥25 |
| 22 | 7.3 | 400±10 | ≥100 | ≥30 |
| 24 | 8.0 | 400±10 | ≥100 | ≥30 |

Table 7: Dimensions of Latex or Red Urethral Catheter

| Size(Fr/Ch) | O.D, ±0.3mm | Overall length, mm |
| --- | --- | --- |
| 6 | 2.0 | 280±10 |
| 8 | 2.7 | 280±10 |
| 10 | 3.3 | 280±10 |
| 12 | 4.0 | 400±10 |
| 14 | 4.7 | 400±10 |
| 16 | 5.3 | 400±10 |
| 18 | 6.0 | 400±10 |
| 20 | 6.7 | 400±10 |
| 22 | 7.3 | 400±10 |
| 24 | 8.0 | 400±10 |

Table 8: Dimensions of Tiemann Urethral Catheter with balloon

| Size(Fr/Ch) | O.D, ±0.3mm | Overall length, mm | Flow rate, ml/min |
| --- | --- | --- | --- |
| 12 | 4.0 | 395±15 | ≥50 |
| 14 | 4.7 | 395±15 | ≥70 |
| 16 | 5.3 | 395±15 | ≥100 |
| 18 | 6.0 | 395±15 | ≥100 |
| 20 | 6.7 | 395±15 | ≥100 |
| 22 | 7.3 | 395±15 | ≥100 |
| 24 | 8.0 | 395±15 | ≥100 |
| 26 | 8.7 | 395±15 | ≥100 |
| 28 | 9.3 | 395±15 | ≥100 |
| 30 | 10.0 | 395±15 | ≥100 |

Table 9: Dimensions of Tiemann Urethral Catheter without balloon

| Size(Fr/Ch) | O.D, ±0.3mm | Overall length, mm | Flow rate, ml/min |
| --- | --- | --- | --- |
| 12 | 4.0 | 395±15 | ≥50 |
| 14 | 4.7 | 395±15 | ≥70 |
| 16 | 5.3 | 395±15 | ≥100 |
| 18 | 6.0 | 395±15 | ≥100 |
| 20 | 6.7 | 395±15 | ≥100 |
| 22 | 7.3 | 395±15 | ≥100 |
| 24 | 8.0 | 395±15 | ≥100 |

4.2.2 Physical tests

Physical tests are performed to the Urethral Catheters in accordance with ENISO20696:2018.

4.3 Chemical & Biological requirements

The Urethral Catheters are sterilized by ETO in accordance with ENISO11135, the index of the biological compatibility conforms to ENISO10993-1.

| Sequence # | Biocompatibility tests performed | Results |
| --- | --- | --- |
| 1 | Cytotoxicity | Comply with ENISO10993-5 |
| 2 | Sensitization | Comply with ISO10993-10 |
| 3 | Intracutaneous reactivity | Comply with ISO10993-10 |
| 4 | EO residue | Comply with ENISO10993-7 |