- Brief introduction

An Indwelling catheter is a tube which is inserted into the bladder and left in place in drain urine.

An indwelling catheter may be inserted into the bladder in two ways:

- Most often, the catheter is inserted through the urethra. This is the tube that carries urine from the bladder to the outside of the body.

- Sometimes, the health care provider will insert a catheter into bladder through a small hole in your belly. This is done at a hospital or health care provider's office.

The Silicone Foley Catheters are available as 2-way Foley Catheters(with balloon, shaft with x-ray line, drainage funnel, inflation funnel, and valve) and 3-way Foley Catheters (with additional irrigation funnel). Also there are two type available: standard type and suprapubic type. The catheter external diameter (Fr/CH) and the balloon inflation volume (ml/cc) are printed on the proximal end of the drainage funnel.

Syringe (optional) can be supplied together with Silicone Foley Catheters. Syringe with sterile aqueous glycerin solution can be used for inflating balloon; The empty syringe can be used to deflate the balloon before the removal of the catheter.

The Silicone Foley Catheters are used in bladder drainage of urine and/or irrigation of fluids by urethra placement They are made of silicone and for short-term use, consist of shaft, drainage funnel, inflation funnel, irrigation funnel (if present), pre-loaded stylet(if present), balloon and valve. The Silicone Foley Catheter has the following features:

- 100% Silicone

- Color-coded for visualization of size

- Alternative types are available

- Sterile by EO

- For single use only

- Hard valves is available

The Silicone Foley Catheters are used during the following situations:

- On patients who are anesthesized or sedated for surgery or other medical care

- On comatose patients

- On some incontinent patients

- On patients whose prostate is enlarged to the point that urine flow from the bladder is cut off. The catheter is kept in until the problem is resolved.

- On patients with acute urinary retention.

- On patients who are unable due to paralysis or physical injury to use either standard toilet facilities or urinals.

- Sometimes before Furosemide administration

- On patients with urethral or pelvic surgery or trauma

Table of Variants

| Type | Standard Type and Suprapubic Type |
| --- | --- |
| Size | 2-way Paediatric | 6Fr, 8 Fr, 10 Fr |
|  | 2-way Adult | 12Fr, 14 Fr, 16 Fr, 18 Fr, 20 Fr, 22 Fr, 24 Fr, 26 Fr |
|  | 3-way Adult | 16 Fr, 18 Fr, 20 Fr, 22 Fr, 24 Fr, 26 Fr |

2. The purpose and functions of the product and its variants and their design

Intended Use: The Silicone Foley Catheter is used in bladder drainage of urine and/or irrigation of fluids by urethra placement. It can be retained in the human body for up to 29 days.

To achieve above-mentioned purpose, the product should have the following functions: draiage of urine and/or inserting fluids into the bladder for medication.

Drainage of urine and medication:

The Silicone Foley Catheter is made of 100% silicone which has appropriate rigidity, it resists to kink and soft enough for urinary catheterization. After inserting, the drainage lumen of the Silicone Foley Catheter shall be able to drain the urine or remove the fluids into the bladder.

3. Classification according to Annex IX of Council Directive 93/42/EEC

The Silicone Foley Catheter is invasive device. It is intended for short-term use. According to rule 5 of Annex IX, III of Council Directive 93/42/EEC (2007), the Silicone Foley Catheter belongs to class IIa medical device.

4. Main technical parameters of the product

4.1 Material

| No | Component | Material used | Spec/Code | Supplier |
| --- | --- | --- | --- | --- |
| 1 | Shaft | Silicone/DCBP | TRK8016-65P | Dong Guan Tongruikai |
| 2 | Opaque line | Silicone | C020R-1523 | Dong Guan Tongruikai |
| 3 | Opaque line ancillary materials | BaSO4 | C020R-1523 | Dong Guan Tongruikai |
| 4 | Balloon | Silicone | SR330A/B | Shenzhen SQUARE SILICONE CO.,LTD |
| 5 | Sleeve | PP | K4912 | KOREA PETROCHEMICALIND.,LTD |
| 6 | Hard valve | PVC | ----- | Chang Zhou Jintan Weiniankang |
| 7 | Tip | Silicone adhesive | --- | Shen Zhen Haiyang Silicone Rubber Material Co., Ltd |
| 8 | Primary Package | Inner pouch | EVA | ----- | Zhangjiagang Hengya |
|  |  | Outer pouch | Coated paper+film | coated paper 65g +0.09mm XPP | Jiaxing Kebang Medical Packing Co.,Ltd |

4.2 Physical requirements

4.2.1 Measures

### Table 1: Dimensions of 2-way

| Size (Fr/Ch) | O.D. (±0.3) mm | Overall length, mm | Color code | Flow rate, ≥ml/min |
| --- | --- | --- | --- | --- |
|  |  |  |  | Drainage lumen |
| 6 | 2.0 | 310±10mm | Pink | 10 |
| 8 | 2.7 |  | Black | 15 |
| 10 | 3.3 |  | Gray | 30 |
| 12 | 4.0 | 400±10mm | White | 50 |
| 14 | 4.7 |  | Green | 70 |
| 16 | 5.3 |  | Orange | 100 |
| 18 | 6.0 |  | Red | 100 |
| 20 | 6.7 |  | Yellow | 100 |
| 22 | 7.3 |  | Purple | 100 |
| 24 | 8.0 |  | Blue | 100 |
| 26 | 8.7 |  | Pink | 100 |

### Table 2: Dimensions of 3-way

| Size (Fr/Ch) | O.D. (±0.3) mm | Overall length, mm | Color code | Flow rate, ≥ml/min |
| --- | --- | --- | --- | --- |
|  |  |  |  | Drainage lumen | Irrigation lumen |
| 16 | 5.3 | 400±10mm | Orange | 100 | 25 |
| 18 | 6.0 |  | Red | 100 | 25 |
| 20 | 6.7 |  | Yellow | 100 | 25 |
| 22 | 7.3 |  | Purple | 100 | 30 |
| 24 | 8.0 |  | Blue | 100 | 30 |
| 26 | 8.7 |  | Pink | 100 | 30 |

Table 3: Balloon capacity

| Size (Fr/Ch) | Balloon capacity (ml), 2-way | Balloon capacity (ml), 3-way |
| --- | --- | --- |
| 6 | 1.5/3 | / |
| 8 | 3-5/5 | / |
| 10 | 3-5/5 | / |
| 12 | 5/10/15/30 | / |
| 14 | 5/10/15/30 | / |
| 16 | 5/10/15/30 | 10/15/30 |
| 18 | 5/10/15/30 | 10/15/30 |
| 20 | 5/10/15/30 | 10/15/30 |
| 22 | 5/10/15/30 | 10/15/30 |
| 24 | 5/10/15/30 | 10/15/30 |
| 26 | 5/10/15/30 | 10/15/30 |

4.2.2 Physical tests

Physical tests are performed to the Silicone Foley Catheters in accordance with ENISO20696:2018

4.3 Chemical & Biological requirements

The Silicone Foley Catheter are sterilized by ETO in accordance with ENISO11135, the index of the biological compatibility conforms to ENISO10993-1.

| No. | Biocompatibility tests performed | Results |
| --- | --- | --- |
| 1 | Cytotoxicity | Comply with ENISO10993-5 |
| 2 | Sensitization | Comply with ISO10993-10 |
| 3 | Intracutaneous reactivity | Comply with ISO10993-10 |
| 4 | EO residue | Comply with ENISO10993-7 |