- Brief introduction

An Endotracheal Tube is used in general anaesthesia, intensive care and emergency medicine for airway management and mechanical ventilation. The tube is inserted into a patient's trachea through the patient's nose or mouth in order to ensure that the airway is not closed off and that air is able to reach the lungs. The Endotracheal Tubess are regarded as the most reliable available method for protecting a patient's airway.

The standard Endotracheal Tube is inserted into a patient's trachea through the patient's nose or mouth in order to maintain an open air passage and prevent aspiration of mucus and gastric contents. It is designed for general anaesthesia, intensive care and emergency case for airway management and mechanical ventilationensure. It is available in cuffed and uncuffed type. The cuffed type consists of tube, cuff, inflation line, valve, pilot balloon and connector; The uncuffed type is consists of tube and connector.

The Endotracheal Tubess Series are made from PVC in medical grade, consists of tube, cuff, inflation line, valve, pilot balloon, connector, bite block(if present), jet tubing(if present). The Endotracheal Tubes has the following features:

- For both oral & nasal intubation

- 100% latex free

- Atraumatic soft rounded beveled tip

- High volume low pressure cuff provides an effective low pressure seal

- Softer rounded Murphy eye is less invasive

- Accurate depth marks

- Standard connector

- Radio opaque line provided

- Blister package or peelable pouch as per request

- Available with DEHP-FREE material as per request

- Sterile by EO, single use

Table of Variants

| Type | Size, I.D(mm) |
| --- | --- |
|  | Cuffed | Uncuffed |
| Oral, Nasal, Standard, Reinforced, Evacuation Lumen (Cuffed, Uncuffed) | I.D. 2.0, I.D. 2.5, I.D.3.0, I.D.3.5, I.D.4.0, I.D.4.5, I.D.5.0, I.D.5.5, I.D.6.0, I.D.6.5, I.D.7.0, I.D.7.5,I.D.8.0, I.D.8.5, I.D.9.0, I.D.9.5, I.D.10.0, ID 10.5, ID 11.0 | I.D. 2.0, I.D. 2.5, I.D.3.0, I.D.3.5, I.D.4.0, I.D.4.5, I.D.5.0, I.D.5.5, I.D.6.0, I.D.6.5, I.D.7.0, I.D.7.5,I.D.8.0, I.D.8.5, I.D.9.0, I.D.9.5, I.D.10.0, ID 10.5, ID 11.0 |

2. The purpose and functions of the product and its variants and their design

Intended Use

- Endotracheal Tubes

Intended Use: An Endotracheal Tubes is used in general anaesthesia, intensive care and emergency medicine for airway management and mechanical ventilation. The tube is inserted into a patient's trachea through the patient's nose or mouth in order to ensure that the airway is not closed off and that air is able to reach the lungs.

To achieve above-mentioned purpose, the product should have the following functions: airway management, mechanical ventilation and oxygen insufflation.

Airway management:

The Endotracheal Tube is made of PVC which has appropriate rigidity, it resists to kink and soft enough for intubating. After intubation, the Endotracheal Tubes shall be able to maintain the airway of patient open.

Mechanical ventilation:

The Endotracheal Tubess have the standard lumen connector which allows the Endotracheal Tubes to be connected with respiratory machine for ventilation.

3. Classification according to Annex IX of Council Directive 93/42/EEC

The Endotracheal Tubess are invasive device. The device is intended for short-term use. According to rule 5 of Annex IX, III of Council Directive 93/42/EEC (2007), the Endotracheal Tubess belong to class Ⅱa medical devices.

4. Main technical parameters of the product

4.1 Material

Material for Endotracheal Tubes as follows:

| No | Component | Material used | Spec/Code | Supplier |
| --- | --- | --- | --- | --- |
| 1 | Main tube | PVC | SME3200ABT7101 | Shenzhen Lite Plastic Co., Ltd |
| 2 | Inflation line | PVC | Lung tube material | Dongguan Xinchen New Material Technology Co., Ltd |
| 3 | Balloons | PVC | SME7200F4BT02 | Shenzhen Lite Plastic Co., Ltd |
| 4 | Pilot balloon | PVC | SME7200F4BT02 | Shenzhen Lite Plastic Co., Ltd |
| 5 | Opaque line | PVC | M-JP3(178RAD4-BL1) | Dongguan Xinchen New Material Technology Co., Ltd |
| 6 | Valve | Main material: PVC | - | Jiangxi zongjie Medical Devices Co., Ltd. |
| 7 | 15mm Connector | PP | - | Jiangxi zongjie Medical Devices Co., Ltd. |
| 8 | tube printing | Pad printing ink | 7502 | Yongyi printing equipment Firm. |
| 9 | spring | Stainless steel | 304 | Ningbo Yuxin Medical Devices Co., Ltd. |
| 10 | Grouting tube | PVC paste resin. | - | Shenzhen Yinochen Materials Co., Ltd. |
| 11 | Primary Package | Inner pouch | Coated paper+film | 60g + PET/PE 45g | Quzhou Zexiao Medical Packaging Material Co., Ltd |
| 12 | medium box | Single-layer corrugated. | 250g whiteboard paper and E-watts. | Guangzhou Chenda Packaging Products Co., Ltd. |
| 13 | Outer box | Five-story double corrugated. | 200g+140g+100g+140g+200g | Zengcheng Weili Paper Packaging Co., Ltd. |

4.2 Physical requirements

4.2.1 Measures

Table 1: Dimensions of Endotracheal Tubes(Oral）

| Size | Minimum Length | ID | OD | Cuff Resting Diameter, mm (±15% of the marked value) | C | S1 |
| --- | --- | --- | --- | --- | --- | --- |
| 2.0 | 130 | 2.0±0.15 | 2.9±0.15 | 9±1.3 | / | ≥70 |
| 2.5 | 140 | 2.5±0.15 | 3.5±0.15 | 9±1.3 | / | ≥70 |
| 3.0 | 160 | 3.0±0.15 | 4.2±0.15 | 11±1.7 | ≤33 | ≥80 |
| 3.5 | 180 | 3.5±0.15 | 4.9±0.15 | 13±2.0 | ≤35 | ≥80 |
| 4.0 | 200 | 4.0±0.15 | 5.5±0.15 | 14±2.1 | ≤41 | ≥90 |
| 4.5 | 220 | 4.5±0.15 | 6.2±0.15 | 14±2.1 | ≤45 | ≥100 |
| 5.0 | 240 | 5.0±0.15 | 6.9±0.15 | 18±2.7 | ≤56 | ≥110 |
| 5.5 | 270 | 5.5±0.15 | 7.5±0.15 | 18±2.7 | ≤56 | ≥120 |
| 6.0 | 280 | 6.0±0.15 | 8.2±0.15 | 22±3.3 | ≤58 | ≥125 |
| 6.5 | 290 | 6.5±0.20 | 8.7±0.20 | 22±3.3 | ≤62 | ≥130 |
| 7.0 | 305 | 7.0±0.20 | 9.3±0.20 | 25±3.7 | ≤66 | ≥135 |
| 7.5 | 310 | 7.5±0.20 | 10.0±0.20 | 25±3.7 | ≤69 | ≥140 |
| 8.0 | 325 | 8.0±0.20 | 10.7±0.20 | 27±4.0 | ≤72 | ≥145 |
| 8.5 | 325 | 8.5±0.20 | 11.3±0.20 | 27±4.0 | ≤75 | ≥155 |
| 9.0 | 325 | 9.0±0.20 | 12.0±0.20 | 29±4.4 | ≤78 | ≥160 |
| 9.5 | 325 | 9.5±0.20 | 12.7±0.20 | 29±4.4 | ≤81 | ≥165 |
| 10.0 | 325 | 10.0±0.20 | 13.3±0.20 | 29±4.4 | ≤85 | ≥170 |
| 10.5 | 325 | 10.5±0.20 | 14.0±0.20 | 31±4.7 | ≤85 | ≥170 |
| 11.0 | 325 | 11.0±0.20 | 14.7±0.20 | 31±4.7 | ≤85 | ≥170 |

Table 2: Dimensions of Endotracheal Tubes(Nasal）

| Size | Minimum Length | ID | OD | Cuff Resting Diameter, mm (±15% of the marked value) | C | S1 |
| --- | --- | --- | --- | --- | --- | --- |
| 2.0 | 130 | 2.0±0.15 | 2.9±0.15 | 9±1.3 | / | ≥70 |
| 2.5 | 140 | 2.5±0.15 | 3.5±0.15 | 9±1.3 | / | ≥70 |
| 3.0 | 160 | 3.0±0.15 | 4.2±0.15 | 11±1.7 | ≤33 | ≥80 |
| 3.5 | 180 | 3.5±0.15 | 4.9±0.15 | 13±2.0 | ≤35 | ≥80 |
| 4.0 | 200 | 4.0±0.15 | 5.5±0.15 | 14±2.1 | ≤41 | ≥90 |
| 4.5 | 220 | 4.5±0.15 | 6.2±0.15 | 14±2.1 | ≤45 | ≥100 |
| 5.0 | 240 | 5.0±0.15 | 6.9±0.15 | 18±2.7 | ≤56 | ≥110 |
| 5.5 | 270 | 5.5±0.15 | 7.5±0.15 | 18±2.7 | ≤56 | ≥120 |
| 6.0 | 280 | 6.0±0.15 | 8.2±0.15 | 22±3.3 | ≤58 | ≥125 |
| 6.5 | 290 | 6.5±0.20 | 8.7±0.20 | 22±3.3 | ≤62 | ≥130 |
| 7.0 | 305 | 7.0±0.20 | 9.3±0.20 | 25±3.7 | ≤66 | ≥135 |
| 7.5 | 310 | 7.5±0.20 | 10.0±0.20 | 25±3.7 | ≤69 | ≥140 |
| 8.0 | 325 | 8.0±0.20 | 10.7±0.20 | 27±4.0 | ≤72 | ≥145 |
| 8.5 | 325 | 8.5±0.20 | 11.3±0.20 | 27±4.0 | ≤75 | ≥155 |
| 9.0 | 325 | 9.0±0.20 | 12.0±0.20 | 29±4.4 | ≤78 | ≥160 |
| 9.5 | 325 | 9.5±0.20 | 12.7±0.20 | 29±4.4 | ≤81 | ≥165 |
| 10.0 | 325 | 10.0±0.20 | 13.3±0.20 | 29±4.4 | ≤85 | ≥170 |
| 10.5 | 325 | 10.5±0.20 | 14.0±0.20 | 31±4.7 | ≤85 | ≥170 |
| 11.0 | 325 | 11.0±0.20 | 14.7±0.20 | 31±4.7 | ≤85 | ≥170 |

Table 3: Dimensions of Endotracheal Tubes(Standard）

| Size | Minimum Length | ID | OD | Cuff Resting Diameter, mm (±15% of the marked value) | C | S1 |
| --- | --- | --- | --- | --- | --- | --- |
| 2.0 | 130 | 2.0±0.15 | 2.9±0.15 | 9±1.3 | / | ≥70 |
| 2.5 | 140 | 2.5±0.15 | 3.5±0.15 | 9±1.3 | / | ≥70 |
| 3.0 | 160 | 3.0±0.15 | 4.2±0.15 | 11±1.7 | ≤33 | ≥80 |
| 3.5 | 180 | 3.5±0.15 | 4.9±0.15 | 13±2.0 | ≤35 | ≥80 |
| 4.0 | 200 | 4.0±0.15 | 5.5±0.15 | 14±2.1 | ≤41 | ≥90 |
| 4.5 | 220 | 4.5±0.15 | 6.2±0.15 | 14±2.1 | ≤45 | ≥100 |
| 5.0 | 240 | 5.0±0.15 | 6.9±0.15 | 18±2.7 | ≤56 | ≥110 |
| 5.5 | 270 | 5.5±0.15 | 7.5±0.15 | 18±2.7 | ≤56 | ≥120 |
| 6.0 | 280 | 6.0±0.15 | 8.2±0.15 | 22±3.3 | ≤58 | ≥125 |
| 6.5 | 290 | 6.5±0.20 | 8.7±0.20 | 22±3.3 | ≤62 | ≥130 |
| 7.0 | 305 | 7.0±0.20 | 9.3±0.20 | 25±3.7 | ≤66 | ≥135 |
| 7.5 | 310 | 7.5±0.20 | 10.0±0.20 | 25±3.7 | ≤69 | ≥140 |
| 8.0 | 325 | 8.0±0.20 | 10.7±0.20 | 27±4.0 | ≤72 | ≥145 |
| 8.5 | 325 | 8.5±0.20 | 11.3±0.20 | 27±4.0 | ≤75 | ≥155 |
| 9.0 | 325 | 9.0±0.20 | 12.0±0.20 | 29±4.4 | ≤78 | ≥160 |
| 9.5 | 325 | 9.5±0.20 | 12.7±0.20 | 29±4.4 | ≤81 | ≥165 |
| 10.0 | 325 | 10.0±0.20 | 13.3±0.20 | 29±4.4 | ≤85 | ≥170 |
| 10.5 | 325 | 10.5±0.20 | 14.0±0.20 | 31±4.7 | ≤85 | ≥170 |
| 11.0 | 325 | 11.0±0.20 | 14.7±0.20 | 31±4.7 | ≤85 | ≥170 |

Table 4: Dimensions of Endotracheal Tubess (Reinforced)

| Size | Minimum Length | ID | OD | Cuff Resting Diameter, mm (±15% of the marked value) | C | S1 |
| --- | --- | --- | --- | --- | --- | --- |
| 3.0 | 160 | 3.0±0.15 | 5.3±0.15 | 11±1.7 | ≤33 | ≥80 |
| 3.5 | 180 | 3.5±0.15 | 5.8±0.15 | 13±2.0 | ≤35 | ≥80 |
| 4.0 | 200 | 4.0±0.15 | 6.3±0.15 | 14±2.1 | ≤41 | ≥90 |
| 4.5 | 220 | 4.5±0.15 | 6.8±0.15 | 14±2.1 | ≤45 | ≥100 |
| 5.0 | 240 | 5.0±0.15 | 7.3±0.15 | 18±2.7 | ≤56 | ≥110 |
| 5.5 | 270 | 5.5±0.15 | 7.8±0.15 | 18±2.7 | ≤56 | ≥120 |
| 6.0 | 280 | 6.0±0.20 | 8.5±0.15 | 22±3.3 | ≤58 | ≥125 |
| 6.5 | 290 | 6.5±0.20 | 9.0±0.20 | 22±3.3 | ≤62 | ≥130 |
| 7.0 | 305 | 7.0±0.20 | 9.6±0.20 | 25±3.7 | ≤66 | ≥135 |
| 7.5 | 310 | 7.5±0.20 | 10.2±0.20 | 25±3.7 | ≤69 | ≥140 |
| 8.0 | 325 | 8.0±0.20 | 10.9±0.20 | 27±4.0 | ≤72 | ≥145 |
| 8.5 | 325 | 8.5±0.20 | 11.5±0.20 | 27±4.0 | ≤75 | ≥155 |
| 9.0 | 325 | 9.0±0.20 | 12.2±0.20 | 29±4.4 | ≤78 | ≥160 |
| 9.5 | 325 | 9.5±0.20 | 12.9±0.20 | 29±4.4 | ≤81 | ≥165 |
| 10.0 | 325 | 10.0±0.20 | 13.5±0.20 | 29±4.4 | ≤85 | ≥170 |
| 10.5 | 325 | 10.5±0.20 | 14.2±0.20 | 31±4.7 | ≤85 | ≥170 |
| 11.0 | 325 | 11.0±0.20 | 14.9±0.20 | 31±4.7 | ≤85 | ≥170 |

Table 5: Dimensions of Endotracheal Tubess (Evacuation Lumen)

| Size | Minimum Length | ID | OD | Cuff Resting Diameter, mm (±15% of the marked value) | C | S1 |
| --- | --- | --- | --- | --- | --- | --- |
| 6.0 | 280 | 6.0±0.15 | 9.0±0.50 | 22±3.3 | ≤58 | ≥125 |
| 6.5 | 290 | 6.5±0.20 | 9.8±0.50 | 22±3.3 | ≤62 | ≥130 |
| 7.0 | 305 | 7.0±0.20 | 10.4±0.50 | 25±3.7 | ≤66 | ≥135 |
| 7.5 | 310 | 7.5±0.20 | 11.2±0.50 | 25±3.7 | ≤69 | ≥140 |
| 8.0 | 325 | 8.0±0.20 | 11.8±0.50 | 27±4.0 | ≤72 | ≥145 |
| 8.5 | 325 | 8.5±0.20 | 12.6±0.50 | 27±4.0 | ≤75 | ≥155 |
| 9.0 | 325 | 9.0±0.20 | 13.1±0.50 | 29±4.4 | ≤78 | ≥160 |

4.2.2 Physical tests

Physical tests are performed to the Endotracheal Tubess in accordance with EN1782:1998+ A1:2009＆ENISO5361:2016

| No. | Test Item | Requirements |
| --- | --- | --- |
| 1 | Surface | The surface shall appear free from extraneous matter |
| 2 | TESTED PRODUCT(I.D.: MM) |
|  | Tube Dimension (mm) | I.D. | Comply with requirement |
|  |  | O.D. |  |
|  |  | Overall length |  |
| 3 | Connector Dimensions (mm) | Length of patient end | Comply with requirement |
|  |  | Length of machine end | ≥15mm |
| 4 | Angle of bevel | 38±10o |
| 5 | Kink resistance | When a tracheal tube is tested for tube kink resistance according to the method described in Annex H of ENISO5361:2016, the steel ball shall pass freely through the lumen of the tracheal tube. |
| 6 | Radius of curvature of tracheal tube (mm) | The Radius of Curvature shall be (140 ±20) mm for tubes of sizes 6,5 and larger;sizes 6,0 and smaller may have a radius of curvature other than (140 ± 20) mm. |
| 7 | Size of the Murphy eye | The area of Murphy eye shall be not less than 80% of the cross-sectional area derived from the minimum inside diameter that size of tube |
|  | Location of the Murphy eye | The location of the murphy eye shall be on the side of the tube opposite the bevel. |
| 8 | Sterility | Comply with EN556-1 |
| 9 | Residual EO | Comply with ENISO10993-7 |
| 10 | Symbols & labeling | Comply with ENISO 15223-1&EN1041 |

4.3 Chemical & Biological requirements

The Endotracheal Tubess are sterilized by ETO in accordance with ENISO11135, the index of the biological compatibility conforms to ENISO10993.2009-1.

| NO. | Biocompatibility tests performed | Results |
| --- | --- | --- |
| 1 | Cytotoxicity | The numerical grade not greater than 2 or the reduction of cell viability not more than 30% |
| 2 | Sensitization | Magnusson and Kligman grades Should less than 1 |
| 3 | Skin Irritation | The mean score should be 1.9 or less |
| 4 | EO residue | EO≤4mg/device, ECH≤9mg/device |