- Brief introduction

Laryngeal Mask Devices are used in anesthesia and in emergency medicine for airway management. They consist of a tube with an inflatable cuff that is inserted into the pharynx. They cause less pain and coughing than an Endotracheal Tube, and are much easier to insert. However, a standard Laryngeal Mask Device does not protect the lungs from aspiration, making them unsuitable for patients at risk for this complication.

The Laryngeal Mask Device is a type of supraglottic airway device that provides an alternative to endotracheal intubation and standard mask anaesthesia in general anaesthesia. It is inserted into the hypopharynx to form a seal around the larynx to allow spontaneous or positive pressure ventilation without the need for intubating the larynx.

The main use of Laryngeal Mask Device is in elective anaesthesia, however it is also useful in situations where manipulation of the head or neck to facilitate ndotracheal intubation is difficult (eg: trapped patient or potential cervical spine injury) or when intubation is unsuccessful. It is not inserted as far as an Endotracheal Tube (it sits in the pharynx, and thus does not need to be inserted into the trachea), and supports both spontaneous and artificial ventilation. It is popular in day case surgery.

- The Silicone Laryngeal Mask Device is made from silicone in medical grade, consists of airway tube, laryngeal mask, connector, Inflating tube, Valve, Pilot balloon, deflation flake(if present), annectent back. The Laryngeal Mask Device has the following features:

Advantages:

- Medical grade silicone for super compatibility and patient comfort

- Allows rapid access

- Does not require laryngoscope

- Relaxants not needed

- Provides airway for spontaneous or controlled ventilation

- Tolerated at lighter anesthetic planes

- 100% latex free

- Peelable pouch or tray available

- Sterile by EO, single use

Disadvantages:

- Does not fully protect against aspiration in the non-fasted patient

- Standard Laryngeal Mask Device does not allow high positive pressure ventilation

- The Disposable PVC Laryngeal Mask Device is made from PVC in medical grade, consists of airway tube, laryngeal mask, connector, Inflating tube, Valve, Pilot balloon, deflation flake (if present), annectant back. The Laryngeal Mask Device has the following features:

Advantages:

- Medical grade PVC with good compatibility

- Allows rapid access

- Does not require laryngoscope

- Relaxants not needed

- Provides airway for spontaneous or controlled ventilation

- Tolerated at lighter anesthetic planes

- 100% latex free

- Peelable pouch or tray available

- Sterile by EO, single use

Disadvantages:

- Does not fully protect against aspiration in the non-fasted patient

- Standard Laryngeal Mask Device does not allow high positive pressure

- The Reusable Reinforced Silicone Laryngeal Mask Device is made from silicone in medical grade, consists of airway tube, laryngeal mask, connector, Inflating tube, Valve, Pilot balloon, deflation flake (if present), annectant back. The Reusable Reinforced Silicone Laryngeal Mask Device has the following features:

Advantages:

- Reinforced airway tube reduces risk of kinking and airway tube occlusion

- Flexible tube allows for placement away from surgical field and improves surgical access while maintaining goog seal

- Medical grade silicone for super compatibility and patient comfort

- Durable high temperature resistant material suitable for autoclave sterilization

- Additional user card for record of device usage times and patient information is supplied in each pouch

- Allows rapid access

- Does not require laryngoscope

- Relaxants not needed

- Provides airway for spontaneous or controlled ventilation

- Tolerated at lighter anesthetic planes

- 100% latex free

- Peelable pouch or tray available

- Sterile by EO, reusable

Disadvantages:

- Does not fully protect against aspiration in the non-fasted patient

- Standard Laryngeal Mask Device does not allow high positive pressure ventilation

- The Disposable Reinforced Silicone Laryngeal Mask Device is made from silicone in medical grade, consists of airway tube, laryngeal mask, connector, Inflating tube, Valve, Pilot balloon, deflation flake(if present), annectent back. The Laryngeal Mask Device has the following features:

Advantages:

- Reinforced airway tube reduces risk of kinking and airway tube occlusion

- Flexible tube allows for placement away from surgical field and improves surgical access while maintaining goog seal

- Medical grade silicone for super compatibility and patient comfort

- Allows rapid access

- Does not require laryngoscope

- Relaxants not needed

- Provides airway for spontaneous or controlled ventilation

- Tolerated at lighter anesthetic planes

- 100% latex free

- Peelable pouch or tray available

- Sterile by EO, single use

Disadvantages:

- Does not fully protect against aspiration in the non-fasted patient

- Standard Laryngeal Mask Device does not allow high positive pressure ventilation

- The Disposable Reinforced PVC Laryngeal Mask Device is made from PVC in medical grade, consists of airway tube, laryngeal mask, connector, Inflating tube, Valve, Pilot balloon, deflation flake(if present), annectent back. The Laryngeal Mask Device has the following features:

Advantages:

- Reinforced airway tube reduces risk of kinking and airway tube occlusion

- Flexible tube allows for placement away from surgical field and improves surgical access while maintaining goog seal.

- Allows rapid access

- Does not require laryngoscope

- Relaxants not needed

- Provides airway for spontaneous or controlled ventilation

- Tolerated at lighter anesthetic planes

- 100% latex free

- Peelable pouch or tray available

- Sterile by EO, single use

Disadvantages:

- Does not fully protect against aspiration in the non-fasted patient

- Standard Laryngeal Mask Device does not allow high positive pressure ventilation

- The Reusable Silicone Laryngeal Mask Device is made from silicone in medical grade, consists of airway tube, laryngeal mask, connector, Inflating tube, Valve, Pilot balloon, deflation flake(if present), annectent back. The Reusable Silicone Laryngeal Mask Device has the following features:

Advantages:

- Medical grade silicone for super compatibility and patient comfort

- Durable high temperature resistant material suitable for autoclave sterilization

- Additional user card for record of device usage times and patient information is supplied in each pouch

- Allows rapid access

- Does not require laryngoscope

- Relaxants not needed

- Provides airway for spontaneous or controlled ventilation

- Tolerated at lighter anesthetic planes

- 100% latex free

- Peelable pouch or tray available

- Sterile by EO, reusable

Disadvantages:

- Does not fully protect against aspiration in the non-fasted patient

- Standard Laryngeal Mask Device does not allow high positive pressure ventilation

Table of Variants

| Type | Sub-Type | Size |
| --- | --- | --- |
| Disposable Silicone Laryngeal Mask Device | One-way Two-way | 1, 1½, 2, 2½, 3, 4, 5 |
| Disposable PVC Laryngeal Mask Device |  | 1, 1½, 2, 2½, 3, 4, 5 |
| Disposable Reinforced PVC Laryngeal Mask Device |  | 1, 1½, 2, 2½, 3, 4, 5 |
| Disposable Reinforced Silicone Laryngeal Mask Device |  | 1, 1½, 2, 2½, 3, 4, 5 |
| Reusable Silicone Laryngeal Mask Device |  | 1, 1½, 2, 2½, 3, 4, 5 |
| Reusable Reinforced Silicone Laryngeal Mask Device |  | 1, 1½, 2, 2½, 3, 4, 5 |

2. The purpose and functions of the product and its variants and their design

Intended Use: Laryngeal Mask Devices are used in anesthesia and in emergency medicine for airway management.

To achieve above-mentioned purpose, the product should have the following functions:

It is inserted blindly into the pharynx, forming a low-pressure seal around the laryngeal inlet and permitting gentle positive pressure ventilation.

3. Classification according to Annex IX of Council Directive 93/42/EEC

The Laryngeal Mask Device is invasive device. The device is intended for short-term use. According to Rule 5 of Annex IX, III of Council Directive 93/42/EEC (2007), the Disposable Silicone Laryngeal Mask Device belongs to class IIa medical devices.

4. Main technical parameters of the product

4.1 Material

Material for Disposable Silicone Laryngeal Mask Device as follows:

| Sequence# | Component | Material used | Spec/Code | Supplier |
| --- | --- | --- | --- | --- |
| 1 | Mask | Silicone | LIM3950 | Shenzhen SenRi |
| 2 | Airway tube | Silicone | C6-265/C6-235/ | Dow Corning Silicone |
| 3 | Inflation tube | Silicone | C6-265/C6-235/ | Dow Corning Silicone |
| 4 | Connector | PC | 2858 | Bayer |
| 5 | Pilot balloon | Silicone | LIM3950 | Shenzhen SenRi |
| 6 | Annectent back | Silicone | LIM3950 | Shenzhen SenRi |
| 7 | Valve | PVC, PP, ABS, Silicone and stainless steel | / | Colorite |
| 8 | Deflation flake | POM | M90-04 | Polyplastics Co., Ltd. |
| 9 | Primary Package | Paper+Film | Uncoated Tyvek +PET/SPE film | Shenzhen Safesecure Co.,Ltd |

Material for Disposable PVC Laryngeal Mask Device as follows:

| Sequence# | Component | Material used | Spec/Code | Supplier |
| --- | --- | --- | --- | --- |
| 1 | Mask | PVC | 7811B-08 | Colorite |
| 2 | Airway tube | PVC | 7811B-08 | Colorite |
| 3 | Inflation tube | PVC | 6511SF140 | Colorite |
| 4 | connector | PC | / | Jiangxi zongjie |
| 5 | Pilot balloon | PVC | 6511SF140 | Colorite |
| 6 | annectent back | PVC | 6511SF140 | Colorite |
| 7 | Valve | PVC, PP, ABS, Silicone and stainless steel | / | Jiangxi zongjie |
| 8 | Deflation flake | POM | / | Guangzhou haopeng |
| 9 | Primary Package | Paper+Film | Uncoated Tyvek +PET/SPE film | Jiangmen Newera Packing Material Co.,LTd. |

Material for Disposable Reinforced Silicone Laryngeal Mask Device as follows:

| Sequence# | Component | Material used | Spec/Code | Supplier |
| --- | --- | --- | --- | --- |
| 1 | Mask | Silicone | RH6250-50 | Guangzhou Ruihe |
| 2 | Airway tube | Silicone | RH7043 | Guangzhou Ruihe |
| 3 | Inflation tube | Silicone | RH7043 | Guangzhou Ruihe |
| 4 | Connector | PC | / | Jiangxi zongjie |
| 5 | Pilot balloon | Silicone | RH6250-50 | Guangzhou Ruihe |
| 6 | Annectent back | Silicone | RH6250-50 | Guangzhou Ruihe |
| 7 | Valve | PVC, PP, ABS, Silicone and stainless steel | / | Jiangxi zongjie |
| 8 | Deflation flake | POM | / | Guangzhou haopeng |
| 9 | Spring | Stainless steel | Stainless steel 304 | Dongguan sijie |
| 10 | Primary Package | Paper+Film | Uncoated Tyvek +PET/SPE film | Jiangmen Newera Packing Material Co.,LTd. |

Material for Disposable Reinforced PVC Laryngeal Mask Device as follows:

| Sequence# | Component | Material used | Spec/Code | Supplier |
| --- | --- | --- | --- | --- |
| 1 | Mask | Silicone | RH6250-50 | Guangzhou Ruihe |
| 2 | Airway tube | Silicone | RH7043 | Guangzhou Ruihe |
| 3 | Inflation tube | Silicone | RH7043 | Guangzhou Ruihe |
| 4 | Connector | PC | / | Jiangxi zongjie |
| 5 | Pilot balloon | Silicone | RH6250-50 | Guangzhou Ruihe |
| 6 | Annectent back | Silicone | RH6250-50 | Guangzhou Ruihe |
| 7 | Valve | PVC, PP, ABS, Silicone and stainless steel | / | Jiangxi zongjie |
| 8 | Deflation flake | POM | / | Guangzhou haopeng |
| 9 | Spring | Stainless steel | Stainless steel 304 | Dong Duan SiJie |
| 10 | Primary Package | Paper+Film | Uncoated Tyvek +PET/SPE film | Jiangmen Newera Packing Material Co.,LTd. |

Material for Reusable Silicone Laryngeal Mask Device as follows:

| Sequence# | Component | Material used | Spec/Code | Supplier |
| --- | --- | --- | --- | --- |
| 1 | Mask | Silicone | RH6250-50 | Guangzhou Ruihe |
| 2 | Airway tube | Silicone | RH7043 | Guangzhou Ruihe |
| 3 | Inflation tube | Silicone | RH7043 | Guangzhou Ruihe |
| 4 | Connector | PSF | / | Jiangxi zongjie |
| 5 | Pilot balloon | Silicone | RH6250-50 | Guangzhou Ruihe |
| 6 | Annectent back | Silicone | RH6250-50 | Guangzhou Ruihe |
| 7 | Valve | PP | / | Jiangxi zongjie |
| 8 | Deflation flake | POM | / | Guangzhou haopeng |
| 9 | Primary Package | Paper+Film | Uncoated Tyvek +PET/SPE film | Jiangmen Newera Packing Material Co.,LTd. |

Material for Reusable Reinforced Silicone Laryngeal Mask Device as follows:

| Sequence# | Component | Material used | Spec/Code | Supplier |
| --- | --- | --- | --- | --- |
| 1 | Silicone | RH6250-50 | Guangzhou Ruihe | Silicone |
| 2 | Silicone | RH7043 | Guangzhou Ruihe | Silicone |
| 3 | Silicone | RH7043 | Guangzhou Ruihe | Silicone |
| 4 | PSF | / | Jiangxi zongjie | PSF |
| 5 | Silicone | RH6250-50 | Guangzhou Ruihe | Silicone |
| 6 | Silicone | RH6250-50 | Guangzhou Ruihe | Silicone |
| 7 | PP | / | Jiangxi zongjie | PP |
| 8 | POM | / | Guangzhou haopeng | POM |
| 9 | Spring | Stainless steel | Stainless steel 304 | Dong Duan SiJie |
| 10 | Primary Package | Paper+Film | Uncoated Tyvek +PET/SPE film | Jiangmen Newera Packing Material Co.,LTd. |

4.2 Physical requirements

4.2.1 Measures

Table 1: Specification for Laryngeal Mask Device

| Size | Maximum cuff inflation volumes, ml | Patient Size |
| --- | --- | --- |
| 1 | 4 | Neonates/Infants up to 5 kg |
| 1½ | 7 | Infants 5 to 10kg |
| 2 | 10 | Infants/children 10 to 20 kg |
| 2½ | 14 | Children 20 to 30 kg |
| 3 | 20 | Children(small adults) 30 to 50 kg |
| 4 | 30 | Adults 50 to 70 kg |
| 5 | 40 | Adults 70 to 100 kg |

Table 2: The Dimensions of Disposable Silicone Laryngeal Mask Device Airway Tube

| Size | I.D, mm | Tolerance | O.D., mm | Tolerance | Length, mm | Tolerance | Radius of curvature, mm | Tolerance |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 5.0 | ±0.3 mm | 8.7 | ±0.3mm | 78 | ±5 mm | 60 | ±30 mm |
| 1½ | 6.0 |  | 10.0 |  | 90 |  | 60 |  |
| 2 | 7.0 |  | 11.3 |  | 111 |  | 80 |  |
| 2½ | 8.5 |  | 13.3 |  | 119 |  | 80 |  |
| 3 | 10.0 |  | 15.3 |  | 161 |  | 100 |  |
| 4 | 10.0 |  | 15.3 |  | 161 |  | 100 |  |
| 5 | 11.0 |  | 16.7 |  | 180 |  | 100 |  |

Table 3: The Dimensions for Disposable Silicone Laryngeal Mask Device Silicone Mask

| Size | Mask length | Tolerance | Mask width | Tolerance |
| --- | --- | --- | --- | --- |
| 1 | 44.5 | ±3.0 | 23.8 | ±3.0 |
| 1½ | 53.0 |  | 29 |  |
| 2 | 61.8 | ±4.0 | 34.2 |  |
| 2½ | 73.6 | ±5.0 | 41.7 |  |
| 3 | 86.8 |  | 49.4 |  |
| 4 | 95.2 |  | 54.3 |  |
| 5 | 106.5 |  | 60.8 | ±4.0 |

Table 3: The Dimensions of Disposable PVC Laryngeal Mask Device Airway Tube

| Size | I.D, mm | Tolerance | O.D, mm | Tolerance | Length, mm | Tolerance | Radius of curvature, mm | Tolerance |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 5.5 | ±0.5mm | 8.3 | ±0.5 mm | 79 | ±5 mm | 100 | ±30 mm |
| 1½ | 6.0 |  | 10.0 |  | 90 |  | 120 |  |
| 2 | 7.0 |  | 11.3 |  | 113 |  | 120 |  |
| 2½ | 9.0 |  | 13.3 |  | 120 |  | 140 |  |
| 3 | 10.0 |  | 15.3 |  | 166 |  | 165 |  |
| 4 | 10.0 |  | 15.3 |  | 166 |  | 165 |  |
| 5 | 11.5 |  | 16.7 |  | 181 |  | 185 |  |

Table 4: The Dimensions for Disposable PVC Laryngeal Mask Device PVC Mask

| Size | Mask length, mm | Tolerance | Mask width, mm | Tolerance |
| --- | --- | --- | --- | --- |
| 1 | 39.5 | ±2.0 | 24.0 | ±2 |
| 1½ | 47.5 |  | 28.5 |  |
| 2 | 56.0 |  | 34.5 |  |
| 2½ | 67.0 |  | 40.0 |  |
| 3 | 79.0 |  | 48.5 |  |
| 4 | 89.0 |  | 54.0 |  |
| 5 | 102.0 |  | 63.0 |  |

Table 5: The Dimensions of Disposable Reinforced PVC Laryngeal Mask Device Airway Tube

| Size | I.D, mm | Tolerance | O.D., mm | Tolerance | Length, mm | Tolerance |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | 5.5 | ±0.5 mm | 8.0 | ±0.5mm | 180 | ±5 mm |
| 1½ | 5.5 |  | 8.0 |  | 180 |  |
| 2 | 6.5 |  | 9.0 |  | 185 |  |
| 2½ | 6.5 |  | 9.0 |  | 185 |  |
| 3 | 8.0 |  | 11.2 |  | 210 |  |
| 4 | 8.0 |  | 11.2 |  | 210 |  |
| 5 | 9.0 |  | 12.7 |  | 240 |  |

Table 6: The Dimensions for Disposable Reinforced PVC Laryngeal Mask Device Mask

| Size | Mask length, mm | Tolerance | Mask width, mm | Tolerance |
| --- | --- | --- | --- | --- |
| 1 | 39.5 | ±2.0 | 24.0 | ±2 |
| 1½ | 47.5 |  | 28.5 |  |
| 2 | 56.0 |  | 34.5 |  |
| 2½ | 67.0 |  | 40.0 |  |
| 3 | 79.0 |  | 48.5 |  |
| 4 | 89.0 |  | 54.0 |  |
| 5 | 102.0 |  | 63.0 |  |

Table7: The Dimensions of Disposable Reinforced Silicone Laryngeal Mask Device Airway Tube

| Size | I.D, mm | Tolerance | O.D., mm | Tolerance | Length, mm | Tolerance |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | 5.5 | ±0.3 mm | 8.7 | ±0.3mm | 170 | ±6 mm |
| 1½ | 5.5 |  | 8.7 |  | 170 |  |
| 2 | 6.5 |  | 9.7 |  | 175 |  |
| 2½ | 6.5 |  | 9.7 |  | 175 |  |
| 3 | 7.2 |  | 11.2 |  | 196 |  |
| 4 | 7.2 |  | 11.2 |  | 196 |  |
| 5 | 8.3 |  | 12.7 |  | 225 |  |

Table 8: The Dimensions for Disposable Reinforced Silicone Laryngeal Mask Device Mask

| Size | Mask length, mm | Tolerance | Mask width, mm | Tolerance |
| --- | --- | --- | --- | --- |
| 1 | 44.5 | ±3.0 | 23.8 | ±3.0 |
| 1½ | 53.0 |  | 29 |  |
| 2 | 61.8 | ±4.0 | 34.2 |  |
| 2½ | 73.6 | ±5.0 | 41.7 |  |
| 3 | 86.8 |  | 49.4 |  |
| 4 | 95.2 |  | 54.3 |  |
| 5 | 106.5 |  | 60.8 | ±4.0 |

Table 9: The Dimensions of Reusable Silicone Laryngeal Mask Device Airway Tube

| Size | I.D, mm | Tolerance | O.D., mm | Tolerance | Length, mm | Tolerance | Radius of curvature, mm | Tolerance |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 5.0 | ±0.3 mm | 8.7 | ±0.3mm | 78 | ±5 mm | 60 | ±30 mm |
| 1½ | 6.0 |  | 10.0 |  | 90 |  | 60 |  |
| 2 | 7.0 |  | 11.3 |  | 111 |  | 80 |  |
| 2½ | 8.5 |  | 13.3 |  | 119 |  | 80 |  |
| 3 | 10.0 |  | 15.3 |  | 161 |  | 100 |  |
| 4 | 10.0 |  | 15.3 |  | 161 |  | 100 |  |
| 5 | 11.0 |  | 16.7 |  | 180 |  | 100 |  |

Table 10: The Dimensions for Reusable Silicone Laryngeal Mask Device Mask

| Size | Mask length | Tolerance | Mask width | Tolerance |
| --- | --- | --- | --- | --- |
| 1 | 44.5 | ±3.0 | 23.8 | ±3.0 |
| 1½ | 53.0 |  | 29 |  |
| 2 | 61.8 | ±4.0 | 34.2 |  |
| 2½ | 73.6 | ±5.0 | 41.7 |  |
| 3 | 86.8 |  | 49.4 |  |
| 4 | 95.2 |  | 54.3 |  |
| 5 | 106.5 |  | 60.8 | ±4.0 |

Table 11: The Dimensions of Reusable Reinforced Silicone Laryngeal Mask Device Airway Tube

| Size | I.D., mm | Tolerance, mm | O.D. ,mm | Tolerance, mm | Length, mm | Tolerance, mm |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | 5.5 | ±0.3 | 8.7 | ±0.3 | 170 | ±6 |
| 1½ | 5.5 |  | 8.7 |  | 170 |  |
| 2 | 6.5 |  | 9.7 |  | 175 |  |
| 2½ | 6.5 |  | 9.7 |  | 175 |  |
| 3 | 7.2 |  | 11.2 |  | 196 |  |
| 4 | 7.2 |  | 11.2 |  | 196 |  |
| 5 | 8.3 |  | 12.7 |  | 225 |  |

Table 12: The Dimensions for Reusable Reinforced Silicone Laryngeal Mask Device Mask

| Size | Mask length, mm | Tolerance, mm | Mask width, mm | Tolerance, mm |
| --- | --- | --- | --- | --- |
| 1 | 44.5 | ±3.0 | 23.8 | ±3.0 |
| 1½ | 53.0 |  | 29 |  |
| 2 | 61.8 | ±4.0 | 34.2 |  |
| 2½ | 73.6 | ±5.0 | 41.7 |  |
| 3 | 86.8 |  | 49.4 |  |
| 4 | 95.2 |  | 54.3 |  |
| 5 | 106.5 |  | 60.8 | ±4.0 |

4.2.2 Physical tests

Physical tests are performed to the Disposable Silicone Laryngeal Mask Device in accordance with ISO11712:2009.

4.3 Chemical & Biological requirements

The Disposable Silicone Laryngeal Mask Device is sterilized by ETO in accordance with ENISO11135, the index of the biological compatibility conforms to ENISO10993-1 .

| Sequence # | Biocompatibility tests performed | Results |
| --- | --- | --- |
| 1 | Cytotoxicity | Comply with ENISO10993-5 |
| 2 | Sensitization | Comply with ISO10993-10 |
| 3 | Intracutaneous reactivity | Comply with ISO10993-10 |
| 4 | EO residue | Comply with ENISO10993-7 |